## DISCLAIMERS AND TERMS OF USE
Sonography Procedures: Transducer Disinfection and Infection Control Worksheet Release Version 2.1 DISCLAIMERS & TERMS OF USE
Sonography Procedures: Transducer Disinfection and Infection Control Worksheet © 2019-2022 Society of Diagnostic Medical Sonography, Plano, Texas. All Rights Reserved.
INTRODUCTION: The Sonography Procedures: Transducer Disinfection and Infection Control Worksheet ("Worksheet") is an educational worksheet used to educate students and sonographers about the analytical process needed before a sonographic procedure to determine the transducer's classification. This helps determine the proper (sterile/non-sterile) transducer cover and coupling agent/gel, and the expected transducer reprocessing. This analytical process is repeated after the procedure to determine if a change to the transducer reprocessing is necessary. The worksheet can also be used by facilities to evaluate their policies and procedures and document variances from the best practice (e.g., due to transducer manufacturer’s IFU limits). There are 2 additional tabs at the bottom of the Worksheet - please read the INSTRUCTIONS before using the SONOGRAPHY PROCEDURES. This Worksheet is based on the Sonographer Best Practices for Infection Prevention and Control: Reprocessing the Ultrasound Transducer ("Best Practices") and is subject to change without notice. The Worksheet is for educational purposes only. NOTE: These best practices do not represent a legal requirement under state or federal law. Consult your facility’s infection control department for additional guidance. Please submit comments and suggestions to executivestaff@sdms.org.
SDMS DISCLAIMER AND TERMS OF USE
DISCLAIMER: THE BEST PRACTICES AND THIS WORKSHEET ARE PROVIDED WITHOUT ANY REPRESENTATIONS OR WARRANTIES, EXPRESS OR IMPLIED. THE AUTHORS AND ALL PARTICIPATING, SUPPORTING, AND ENDORSING ORGANIZATIONS EXPRESSLY DISCLAIM ALL LIABILITY TO ANY PARTY FOR THE ACCURACY, COMPLETENESS, OR AVAILABILITY OF THE BEST PRACTICES OR THIS WORKSHEET, OR FOR DAMAGES ARISING OUT OF THE USE OF THE BEST PRACTICES OR WORKSHEET AND ANY INFORMATION CONTAINED WITHIN. NOTE: These Best Practices do not represent a legal requirement under state or federal law. Consult your facility’s infection control department for additional guidance. | .

## INSTRUCTIONS
Sonography Procedures: Transducer Disinfection and Infection Control Worksheet Release Version 2.1 INSTRUCTIONS
WORKSHEET COLUMNS
Columns with a number/circle in BLACK (e.g., ③) indicate the user may make an entry in the column. While the existing entries in the Worksheet use an X to indicate the assessed level of risk for the procedure, any letter, number, or symbol may be used. Columns with a number/circle in RED (e.g., ①) do not allow user data-entry; however, columns also marked with the 🄰 symbol (e.g., ①🄰 and ②🄰 ) indicate the user can add sonography procedures, but only in the provided blank rows at the bottom of the Worksheet.
WORKSHEET ROWS
Each sonography procedure is listed in a Worksheet row. Procedures are grouped in categories. Additional space is provided at the bottom of the Worksheet to allow the user to add up to 20 additional sonography procedures. The user can also enter their Assessed: Before Procedure and Actual: During/After Procedure risk assessment(s) for the added procedures (and the Calculated Classification and Best Practices sections will automatically be determined in the same manner as the sonography procedures provided in the Best Practices). EXISTING VS USER ENTERED RISK ASSESSMENTS: While the letter "X" has been used in the Worksheet to provide guidance on a listed procedure's assessed or actual risk, please remember that anything (e.g., space, number, or character) intentionally or accidentally entered in the Assessed: Before Procedure section (③④⑤) or an Actual: During/After Procedure section (⑥⑦⑧) could impact the Calculated Classification and Best Practices. At least one entry in the Assessed: Before Procedure section (③④⑤) or the Actual: During/After Procedure section (⑥⑦⑧) is needed to properly determine the Calculated Classfication and Best Practices. Also, be sure to only have one entry in an Assessed: Before Procedure section (③④⑤) or an Actual: During/After Procedure section (⑥⑦⑧).
NOTES
1. The Worksheet will use the highest level of risk from the Assessed: Before Procedure section. If the user places an X in both the green (Intact Skin) ③ and the red (Sterile Tissue or Blood) ⑤ columns, only the red column will be used to determine the Calculated Classification and Best Practices. 2. Likewise, the Worksheet will use the highest level of risk from the Actual: During/After Procedure section. If the user places an X in both the green (Intact Skin) ⑥ and the red (Sterile Tissue or Blood) ⑧ columns, only the red column will be used to determine the Calculated Classification and Best Practices. The Actual: During/After Procedure risk assessment will always override a procedure's Assessed: Before Procedure.
SONOGRAPHY PROCEDURES
The Procedure Category ① and Procedure Description ② are provided for more than 70 sonography procedures.
ASSESS EXPECTED AND ACTUAL RISK
The user can edit the risk assessment for the Assessed: Before Procedure section (③④⑤) or Actual: During/After Procedure section (⑥⑦⑧). The Calculated Classification and Best Practices will be automatically updated if the user places any number or character in these columns. NOTE: If the user places something in more than one column in a row (e.g., "X" in both Intact Skin and Mucous Membrane or Non-Intact Skin), only the higher risk (i.e., Mucous Membrane or Non-Intact Skin) will be used to determine the Calculated Classification and Best Practice. Use of the Actual: During/After Procedure section is optional and demonstrates the impact of events that occur during or after a procedure (e.g., transducer cover tear, contact with sterile tissue/blood). If the user places an "X" (or any other number or character) in one of the Actual: During/After Procedure section (⑥⑦⑧), this entry will override the Assessed: Before Procedure (③④⑤) entry, if it is a higher risk. Because the expected risk before the procedure is used to determine the classification, a lower risk entry in the Actual: During/After Procedure section (⑥⑦⑧) will be ignored. EXAMPLE: If Expected: Before Procedure was Intact Skin ③, but the Actual: During/After Procedure was Sterile Tissue or Blood ⑧, the Calculated Classification will change to Critical (and the Best Practices updated accordingly). This change could impact the disinfection procedures for the transducer in this example.
SPAULDING CLASSIFICATION & BEST PRACTICES
The Worksheet determines the Calculated Classification (⑨) and Best Practices (⑩⑪⑫⑬⑭⑮⑯) based on the entries in the Assessed: Before Procedure (③④⑤) and/or Actual: During/After Procedure section (⑥⑦⑧). While the Worksheet's default entry ("X") for each procedure is provided, the user can change the entry, if needed.
COMPLIANCE DOCUMENTATION
This column (⑰) is optional, but can be used to document the facility's variations, if any, from the Calculated Classification and Best Practices (e.g., due to compliance with a manufacturer's Instructions for Use or IFU). An area is provided at the bottom of the Worksheet for any additional user notes.
NOTES
Additional space is provided to enter notes at the bottom of the table.
COMMENTS & SUGGESTIONS
Please submit comments or suggestions to executivestaff@sdms.org.
Sonography Procedures: Transducer Disinfection and Infection Control Worksheet © 2019-2022 Society of Diagnostic Medical Sonography, Plano, Texas. All Rights Reserved.

## SONOGRAPHY PROCEDURES
Sonography Procedures: Transducer Disinfection and Infection Control Worksheet v2.1
The list of sonography procedures below describes the transducer classification and best practices. Note that the assessed classification for a sonography procedure could change based on patient condition (e.g., has open wound or infection in the scan area) or during the procedure (e.g., blood is present). Follow the same assessment of risk and classification for sonography procedures not listed.
EXAMPLE SONOGRAPHY PROCEDURES |  | ASSESSED: BEFORE PROCEDURE |  |  | ACTUAL: DURING/AFTER PROCEDURE |  |  |  | KEY: ✔️* = Best Practice; ✔️ = Allowed; 🚫 = Not Recommended (unless transducer cannot be sterilized or HLD processed)
PROCEDURE CATEGORY | PROCEDURE DESCRIPTION | ASSESS RISK THE TRANSDUCER WILL COME IN CONTACT WITH: |  |  | ASSESS RISK THE TRANSDUCER CAME IN CONTACT WITH: |  |  | CALCULATED CLASSIFICATION | BEST PRACTICES | [6 empty] | FACILITY COMPLIANCE WITH BEST PRACTICE FOR CALCULATED CLASSIFICATION
[9 empty] | TRANSDUCER REPROCCESSING |  |  | TRANSDUCER COVER |  | COUPLING AGENT/GEL
①🄰 | ②🄰 | Intact Skin ③ | Mucous Membrane or Non-Intact Skin ④ | Sterile Tissue or Blood ⑤ | Intact Skin ⑥ | Mucous Membrane or Non-Intact Skin ⑦ | Sterile Tissue or Blood ⑧ | ⑨ | Clean and Sterile Process ⑩ | Clean and HLD Process ⑪ | Clean and LLD Process ⑫ | Sterile ⑬ | Non-Sterile ⑭ | Sterile ⑮ | Non-Sterile ⑯ | Identify and document variance(s) from the calculated classification and best practices, the reason/rationale for each variance(s), and any corrective action(s) to be taken (if applicable) ⑰
Head/Neck | Neck, Thyroid/Parathyroid | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Head/Neck | Neonatal Brain | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Head/Neck | Ophthalmic |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Head/Neck | Spinal Canal & Contents | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Chest/Breast | Breast w/ axilla | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Chest/Breast | Chest (includes mediastinum, chest wall, and upper back) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Extremities | Infant Hips | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Extremities | Non-Vascular Extremity | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Abdomen | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Abdomen Elastography | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Aorta (AAA screening) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Renal Retroperitoneal | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Retroperitoneal - Transplanted Kidney w/ Duplex Doppler | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Scrotum & Testicles | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Abdomen | Transrectal & Transrectal Prostate |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Non-Obstetrical Pelvic | Pelvic (non-OB, transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Non-Obstetrical Pelvic | Pelvic (non-OB, transvaginal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Non-Obstetrical Pelvic | Sonophysterography w/ Doppler (non-OB, transvaginal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Obstetric | Fetal Biophysical Profile | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Fetal Middle Cerebral Artery | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Fetal Umbilical Artery | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant < 14 weeks (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant > 14 weeks (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant (transvaginal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Obstetric | Pregnant w/ Detailed Fetal Anatomic Exam (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant w/ First Trim Fetal Nuchal Translucency (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Obstetric | Pregnant w/ First Trim Fetal Nuchal Translucency (transvaginal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Echocardiography (Fetal) | Fetal Doppler Echocardiography | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Echocardiography (Fetal) | Fetal Echocardiography (2D) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Echocardiography (Pediatric/Adult) | Doppler Echocardiography | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Echocardiography (Pediatric/Adult) | Echocardiography (transesophageal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Echocardiography (Pediatric/Adult) | Echocardiography (transthoracic) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Echocardiography (Pediatric/Adult) | Stress Echocardiography | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Abdominal Duplex Arterial/Venous (transabdominal) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Abdominal Duplex Arterial/Venous (transvaginal) |  | X |  |  |  |  | Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️
Vascular | Carotid Doppler | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Carotid Intima-Media Thickness | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Aorta, IVC, iliac, or Bypass Grafts | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Arterial/Venous Penile | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Extracranial Arteries (Doppler) | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Hemodialysis Access | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Lower/Upper Extremity Arterial | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Duplex Lower/Upper Extremity Veins | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Vascular | Transcranial Doppler Intracranial Artery | X | [5 empty] | Non-Critical | optional | optional | ✔️* | optional | optional | optional | ✔️*
Other | Intraoperative |  |  | X |  |  |  | Critical | ✔️* | ✔️ | 🚫 | ✔️* | 🚫 | ✔️* | 🚫
🄰 = No data-entry permitted in this column except in blank/empty rows at the bottom of worksheet.
USER/FACILITY NOTES:

## CATEGORIES (hidden)
Select Category |  | OPTIONS
X
Abdomen
Chest/Breast
Echocardiography (Fetal)
Echocardiography (Pediatric/Adult)
Extremities
Guidance
Head/Neck
Non-Obstetrical Pelvic
Obstetric
Other
Vascular
NOTES:
1 | A blank row is provided at the top of the list so user can de-select all options from the dropdown list.
2 | If additional categories are desired, they should be added by expanding the table or inserting a new row within the table. Do not simply add at the end without making it part of the table.

## LOOKUPTABLE (hidden)
CLASSIFICATION | PROCESS-STERILE | PROCESS-HLD | PROCESS-LLD | COVER-STERILE | COVER-NON-STERILE | GEL-STERILE | GEL-CLEAN | NOTES
Critical | ✔️* | ✔️ | 🚫 | ✔️* | 🚫 | ✔️* | 🚫 | Disinfection: Sterile (preferred)
Semi-critical | optional | ✔️* | 🚫 | ✔️* | ✔️ | ✔️* | ✔️ | Gel: Single-use sterile (preferred)
Non-critical | optional | optional | ✔️* | optional | optional | optional | ✔️* | Gel: Single-use clean (preferred)
Non-critical (Guidance ¹) | optional | ✔️* | ✔️ | ✔️* | ✔️ | ✔️* | 🚫
KEY: ✔️* = Allowed/Preferred; ✔️ = Allowed; 🚫 = Not Allowed
NOTES:
This lookup table is used in the Sonography Procedures worksheet to automatically generate the Calculated Spaulding Classification and Best Practices columns.
The formulas in the Sonography Procedures worksheet expect the Classifications (column A) exactly as provided above. If changed here, all the applicable formulas must change there!

## DOCUMENTATION (hidden)
DOCUMENTATION
EXAMPLE FORMULA
=IF($A80="Guidance",IF(OR($E80<>"",$H80<>""),"Critical",IF(OR($D80<>"",$G80<>""),"Semi-critical",IF(AND(OR($C80<>"",$F80<>""),$A80="Guidance"),"Non-Critical (Guidance ¹)",""))),IF(OR($E80<>"",$H80<>""),"Critical",IF(OR($D80<>"",$G80<>""),"Semi-critical",IF(OR($C80<>"",$F80<>""),"Non-Critical",""))))
The first half of the formula determines whether "Guidance" is listed in the "Procedure Category" column. If so, the first set of formulas apply - this results in "Non-Critical (Guidance ¹)" being displayed if there is text either of the the "Intact Skin" (③⑥) columns for the row. If any other "Procedure Category" is listed (or selected by the user for rows at the bottom of the Worksheet), the second set of formulas apply and will display Non-Critical, Semi-Critical, or Critical as the Calculated Spaulding Classification (based solely on the entries in the Expected: Before Procedure columns (③④⑤) or Actual: During/After Procedure columns (⑥⑦⑧).
FORMULA LOGIC SUMMARY
IF "Guidance" appears in $A
IF $E OR $H contain something (e.g., "X") Display "Critical"
OR
IF $D OR $G contain something (e.g., "X") Display "Semi-Critical"
OR
IF $C OR $F contain something (e.g., "X") AND $A is "Guidance" Display "Non-Critical (Guidance ¹)" OR Display "Non-Critical"
OR
IF $E OR $H contain something (e.g., "X") Display "Critical" OR IF $D OR $G contain something (e.g., "X") Display "Semi-Critical" OR IF $C OR $F contain something (e.g., "X") Display "Non-Critical"
OR
If none of the criteria above apply, just display and empty cell
PROCESS FOR CREATION OF RESTRICTED VERSION FOR RELEASE
Note: Save all changes to the master file before creating a restricted version for release
1. Turn off Headings (column/row) in each of the first 3 worksheet: View > Headings (uncheck)
2. Turn off Formula Bar in each of the first 3 worksheets: View > Forumla Bar (uncheck)
3. Password protect all sheets individually
a. Deselect all options except "Select unlocked cells"
a. Use the same password for all sheets (and Workbook)
4. Hide the CATEGORIES, LOOKUPTABLE, and DOCUMENTATION tabs (yellow)
5. Select the first cell in the first worksheet (DISCLAIMERS AND TERMS OF USE)
6. Protect the workbook (using same password as worksheets)
7. Save using "Save As…" and name the file with the version number